Gaussian curve fit analysis of Tm (GcTm)

Readme and FAQ

Christoffer Nellåker

Current version: v1.01a

Contents:

Introduction

Installation

How to use

FAQ

Version history

Known bugs

Introduction

This program is designed to analyze dissociation curve data from an ABI Prism 7000 with the SDS v1.2.3. For information as to it use and implications of this please read

(α) Molecular beacon based temperature control and automated analyses for improved resolution of melting temperature analysis using SYBR I Green chemistry – Improved melting temperature determination. Christoffer Nellåker, Ulf Wållgren and Håkan Karlsson. Clinical Chemistry (accepted 2006-10-19).

Also please refer to this in any publications using GcTm or derivations thereof.

Use this program at your own risk.

Installation

1. Unpack the files into a suitable directory.

2. You must have a licensed copy of MATLAB™ (The MathWorks).

3. Either browse to the directory every time you use it or set it in the MATLAB File→ Set Path…

How to use

SDS software:

Pre-PCR checklist:

- Make sure the dye is named SYBR in the SDS software prior to start.

- If a Tmprobe, as described in (α), is to be used make sure the dye is named TAMRA.

- Add a dissociation step.

As described in (α), if a Tm-probe is used the dissociation step should be run three times! After the PCR has completed and one dissociation step has been run start a new dissocitaion protocol and once more after that.

Preparing the data:

- Visually go through the derivative dissociation curves for each well. Wells with no clear peak in the SYBR detector should be removed by checking the “Omit Well” box. Usually these will be handled correctly by the program but the occasional one will cause problems.

- File → Export → Dissociation → Raw&Derivative Data…

Gaussian curve fit analysis of Tm:

Using GcTm:

- Start MATLAB and type GcTm in the command window (make sure you are in the correct directory or that the directory has been set in MATLAB File→ Set Path…)

- Click “File” and find your *.csv data file generated by the SDS program.

- To only run the selected file click “Run”. To run all data files in the selected directory click “Run All”.

Interpreting results

Check the data and verify that it seems reasonable. If When errors are made, this is where they can be detected.

- In the GcTm program window the data can be visualized.

  - “Plot Mesh” will plot the calculated Tms of the SYBR amplicons over the 96-well plate. If Tm-probe data was available a mesh showing the SYBR Tm minus the corresponding TAMRA Tm-probe Tm plus the average of all the Tm-probe Tms used on that 96-well plate (corrected Tm).

  - “Plot well” will plot the data for the well number below, SYBR and TAMRA if available. These plots show the derivative dissociation data, the data range used for curve fitting and the curve fitted to the data. It is recommended to check all wells for errors with these plots, it will be most apparent here.

  - The graph window that is always present is the 96-wells represented in a line for SYBR and the corrected Tm allowing deviations to be detected easily

- GcTm generates an Excel results file at the end of each run. The file is named as your original exported file but with an R appended on the end. This excel file is constructed as follows;

| Detector | Well nr | Tm | Corrected Tm | a → e | n1→ n6 |
| --- | --- | --- | --- | --- | --- |
| 1 (=SYBR) 2 (=TAMRA) | 1-96 | Calculated Tm from the data | If probe was used this will show the corrected Tm (only in detector 2) | Variables describing the curve fitted to the data | The points describing the data ranges used for curve fitting |

3. If triple dissociations have been run, average the corrected Tms for each well over the three dissociations.

FAQ

Q. When I use the Tmprobe the corrected Tms vary more than the uncorrected SYBR Tms.

A. Check the concentration of the Tmprobe you are using, you may have to use more depending on how much SYBR fluorescence there is. Check that evaporation effects have not altered sample concentrations during the PCR

Q. GcTm freezes during calculations and will not respond.

A. Most likely some obscure data has halted a calculation step. Go beck into the SDS program and check your wells. Omit any suspicious wells and try again. If problem persists, systematically analyse half the plate and search by trail and error for the problematic well (Tip: the well on which the calculation halted is ±5 wells from the error). Suspicious data can be data with; infinite points, large amounts of noise, spurious peaks (positive or negative)…

Q. I can not detect the Tmprobe melting.

A. Check the concentration of Tmprobe you are using. The concentration described in (α) is the minimal concentration using a fresh probe. The shelf-life of these types of probes are usually quite short and the concentration you use must compensate for this. As a rule of thumb; If you can’t see the dip in the negative negative derivative curve around 85°C then neither will GcTm. The example data files included show the minimal change required.

Version history

v1.01a

-1st alpha version.

Known bugs

  - Negative negative derivative data can cause a crash.

  - Erroneous and double peaks can cause crash